## RiskWorkSheet
Risk Assessment Form | [5 empty] | Severity | Probablity | Exposure |  |  |  | Values | Risk Level | Action
Group Name: | Group 17 |  |  | Members: |  | Catastrophic - 5 | Probable - 5 | Great - 4 |  | Need mitigation |  | 80-100 | Very High | Discontinue/STOP
|  |  |  | John Ostrander |  | Major - 4 | > 50% - 4 | Above Avg - 3 |  | Acceptable, but could be better |  | 60-79 | High | Immediate Correction
|  |  |  | John Brinkman |  | Serious - 3 | 50% - 3 | Average - 2 |  | Acceptable |  | 40-59 | Substantial | Correction Required
|  |  |  | Brian Schutter |  | Minimal - 2 | Unlikely - 2 | Low or None - 1 |  |  |  | 20-39 | Possible | Attention Needed
[6 empty] | Slight - 1 | Remote - 1 |  |  |  |  | 2019-01-19 | Slight | Possibly Acceptible
Use Risk caclulation sheet to calculate risk.
Item | User | Task | Hazard Category | Hazard | Cause/Failure Mode | Severity | Probability | Exposure | Risk Level | Method of Risk Reduction/Mitigation | Severity | Probability | Exposure | Risk Level | Person Responsible | Date | Comments
Transducer | Nurse | Pressure Measurement |  | Loss of power | Arduino becomes unplugged | 4 | 3 | 3 | 36 | LED light on device to indicate power | 4 | 2 | 3 | 24 | John Brinkman | 2020-04-19
Alternator | Fabricator | Pressure Measurement |  | Incorrect pressure measurement | Valve alternator does not reach the open position | 5 | 2 | 3 | 30 | Design the mechanism to limit movement between open an closed | 5 | 1 | 3 | 15 | Brian Schutter | 2020-04-19
Transducer | Nurse | Pressure Measurement |  | Incorrect pressure measurement | failure to calibrate correctly | 5 | 2 | 1 | 10 | Require confirmation by nurse that callibration is completed | 5 | 1 | 1 | 5 | John Brinkman | 2020-04-19
Transducer | Nurse | Pressure Measurement |  | Incorrect pressure measurement | transducer connected into the drainage system incorrectly | 5 | 2 | 1 | 10 | User manual | 5 | 1 | 1 | 5 | John Ostrander | 2020-04-19
Alternator | Patient | Flow Control |  | Overdrainage | Patient movement | 5 | 4 | 3 | 60 | Accelerometer on patient signals stopcock to close drain | 3 | 2 | 3 | 18 | John Brinkman | 2020-04-19
Lift | Fabricator | Pressure Control |  | Over/Underdrainage | Bag missaligned | 5 | 3 | 3 | 45 | calibration software interacts with nursing staff | 5 | 1 | 3 | 15 | John Brinkman | 2020-04-19
Alternator | Nurse | Pressure Control |  | Overdrainage | Patient being transferred around the hospital while draining | 5 | 4 | 3 | 60 | manual setting to turn the alternator to closed | 5 | 1 | 3 | 15 | John Brinkman | 2020-04-19
Wiring | Fabricator | Pressure Measurement |  | Transducer to Arduino Disconnect | Wiring comes loose on bread board | 5 | 2 | 1 | 10 | Potting the wires in place | 5 | 1 | 1 | 5 | Brian Schutter | 2020-04-19
Wiring | Fabricator | Pressure Control |  | Arduino to Control system disconnect | Wiring comes loose on bread board | 5 | 2 | 1 | 10 | Potting the wires in place | 5 | 1 | 1 | 5 | Brian Schutter | 2020-04-19
Cable | Nurse | Pressure Control |  | Drainage bag falls | Cable breaks | 5 | 2 | 3 | 30 | Use a braided nylon cord | 5 | 1 | 3 | 15 | John Ostrander | 2020-04-19
Wiring | All Users | Insulation |  | Electrical shock | Loose/exposed wires | 2 | 2 | 2 | 8 | manufacturers manual require a check for exposed wires | 2 | 1 | 1 | 2 | John Ostrander | 2020-04-19
Lift | Nurse | Pressure Control |  | Over/Underdrainage | Bag not connected to the correct lift stand | 5 | 2 | 1 | 10 | User manual | 5 | 1 | 1 | 5 | John Ostrander | 2020-04-19
Lift | Nurse | Pressure Control |  | Drainage bag falls | Cable slips off the wheel | 5 | 2 | 3 | 30 | Cable is attached to the wheel and inside the wheel is treated with an adhesive | 5 | 1 | 1 | 5 | John Ostrander | 2020-04-19
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Lists
Severity | Exposure | Probability | Risk | Users | Hazard Category | Hazard
Catastrophic | Frequent | Probable | High | All Users | Biological/Health
Serious | Occasional | Possible | Medium | Doctor | Mechanical
Slight | Remote | Unlikely | Low | Nurse | Electrical
Minimal | None | Negligible |  | Technician | Slip/Trip/Fall
|  |  |  | Fabricator | Eronomic/Human Factors
|  |  |  | Housekeeping | Fire/Explosion
|  |  |  | Paramedic | Heat/Temperature
|  |  |  | Nurse Aide | Noise/Vibration
|  |  |  | Service Technician | Ingress/Egress
|  |  |  | Patient | Material Handling
|  |  |  | Family | Environmental/Industrial Hygene
|  |  |  | Public | Ventilation/Confined Space
Chemicals/Gases
Fluid Pressure
Radiation
Laser
Wastes