## Slide 1: External Ventricular Drain (EVD)
Group 17
John Brinkman, John Ostrander, Brian Schutter
Notes: Recap – Hydrocephalus – too much intracranialpressure (ICP) so need to drain cerebrospinal fluid from the brain

## Slide 2: Specification Update
Specification | Metric | Value
Catheter | Material | Biocompatible, non-allergenic
Fluid Data Acquisition | Flow rate | 100.0 Hz sampling rate
 | Ion concentrations in CSF | 100.0 Hz sampling rate
Drainage Control | Flow rate control | 0.30-0.35 mL/min
 | External pressure control | 2.50 - 3.00 kPa
Patient Monitoring | General Movement | 10.0 Hz sampling rate
 | Head movement and position | 10.0 Hz sampling rate
Cost | Dollars | Sum<$200
Notes: Basically kept everything the same except narrowed the value windows, fixed the sampling rate to accommodate pulsing, and added cost specification. Didn’t change anything else

## Slide 3: Overview of Project
2 basic problems to solve
  Measuring Intracranial Pressure (ICP)
  Controlling ICP / drainage flow
Separated designs into these 2 components
  Try to find best combination
    Combine with microcontroller (Arduino)
  (Except one design that does both)
Note:
  CSF: Cerebrospinal fluid
  ICP: Intracranial pressure
Notes: Use Arduino to connect the 2

## Slide 4: Measurement Designs
Arduino Pressure Sensor
PSI Gauge Pressure Sensor
Flow Rate Pressure Monitor
Blood Pressure Catheter
CSF collection calculation
Simultaneous Pressure Controller and Sensor

## Slide 5: Arduino Pressure Sensor
Pros:
  Cheap
  Designed for Arduino
  In series with catheter
Cons
  High flow range: (1 – 60 L/min) vs .2 mL/min CSF flow
  Low accuracy
Notes: Small wheel inside with magnet attached - Measure flow then calculate pressure

## Slide 6: PSI Gauge Pressure Sensor
Pros
  Arduino compatible
  Accurate
  Correct range
Cons
  Unclear data format
  Minimum Quantity is 50
    Expensive
Notes: Parallel Connection with drainage – static measurement
Piezoelectric technology to measure pressure - accurate
Piezoelectric Effect is the ability of certain materials to generate an electric charge in response to applied mechanical stress – Directly connects to bread board

## Slide 7: Flow Rate Pressure Monitor
Pros
  Can be connected in series with flow
  Measures flow pressure
Cons
  Requires PowerMAT data processor
  Measure high flows and pressures
Notes: Similar to Arduino flow sensor, but more sophisticated
powerMAT increases cost

## Slide 8: Blood Pressure Catheter
Pros
  Adapting already widely used device
  Accurate
  Correct range (~2 kPa)
Cons
  Doesn’t connect to Arduino directly
  Costly
Notes: Sensing catheter inserted into guide catheter – Another concern - When using for blood pressure, not also draining it

## Slide 9: CSF Collection Calculation
Pros
  Simple Design and calculation
  Easy to integrate into current system & to Arduino
  Cheap
Cons
  Inaccurate
  Measures flow and pressure indirectly
Notes: Place force sensor between bag and stand, measure weight difference over time and convert that to flow rate

## Slide 10: Simultaneous Pressure Controller and Sensor
Pros
  Pressure is sensed and controlled simultaneously
Cons
  Not directly compatible with Arduino
  Expensive
  Used for industry
    Not catheter compatible
    Designed for higher flows
Notes: Diaphragm type piezoresistive sensor ----

## Slide 11
Sensors Pugh Chart |  |  |  |  |  |  | 
 | Weight | Design # 1: Arduino Pressure Sensor | Design # 2: PSI Gauge | Design # 3: Flow Pressure Sensor | Design # 4: Blood Pressure Catheter | Design # 5: CSF Collection Calculation | Design # 6: Simultaneous Pressure Sensor/Controller
Weight | 5 | 7 | 4 | 7 | 8 | 5 | 7
Volume | 3 | 3 | 8 | 8 | 9 | 7 | 5
Catheter Diameter | 6 | 6 | 7 | 8 | 10 | 10 | 5
Catheter Length | 4 | 5 | 5 | 6 | 10 | 10 | 3
Material | 3 | 6 | 6 | 8 | 10 | 10 | 4
Compatibility Time | 3 | 8 | 8 | 8 | 9 | 10 | 5
Portability | 2 | 8 | 8 | 6 | 9 | 9 | 8
Interface (data output) | 8 | 10 | 7 | 8 | 7 | 8 | 10
Fluid Data Acquisition | 10 | 6 | 7 | 7 | 9 | 2 | 10
Ion Concentration Measurement (Y/N) | 1 | 0 | 0 | 0 | 0 | 0 | 0
Pressure Measurement Error | 7 | 5 | 8 | 6 | 8 | 5 | 8
Cost (higher=cheaper) | 7 | 8 | 5 | 6 | 3 | 9 | 2
Operational Pressure/Flow Range | 10 | 3 | 8 | 3 | 8 | 2 | 7
Power Requirement | 6 | 4 | 6 | 8 | 5 | 8 | 4
Score |  | 443 | 497 | 487 | 575 | 474 | 479

## Slide 12: Final Sensing Decision
Blood pressure Catheter
  Chosen for
    Best accuracy and measurement range
    Handles pulsing
  Next step
    Connect with Arduino
    Ensure draining and sensing can work simultaneously

## Slide 13: Controlling Designs
Simultaneous pressure controller and sensor
Current Method – manual drainage bag height adjustment
I/Q valve flow rate controller
Automated drainage bag height adjuster

## Slide 14: Manual drainage bag height adjustment
Pros
  Currently used
  Relatively reliable
Cons
  Requires constant care in the ICU
  Patients cannot move
  negligible improvement to current system
Notes: Flow rate and pressure determined by height of bag
Height determined by neurosurgeon
Only change compared to before would be an alert for dangerous flow/pressure

## Slide 15: I/Q valve flow rate controller
Pros
  In series
  Relatively simple
Cons
  Intended for industry use – adapter necessary
  Accuracy decreases with smaller flow
Notes: Electrical signal modulates flow through apparatus – just opens / closes orifice

## Slide 16: Automated Drainage Bag Height Adjuster
Pros
  Simple
  Cheap
  Basic principles already in use
  Automated
Cons
  Slow response time
  Adds some bulk
Notes: Arduino motor + pulley system to control height
Add shut off valve to bag or catheter for motor movements and emergency situations to increase patient safety

## Slide 17
Controllers Pugh Chart |  |  |  |  | 
 | Weight | Design # 6: Simultaneous Pressure Sensor/Controller | Design # 7: Leave it Manual | Design # 8 : I/Q Valve | Design # 9: Automated Drainage Bag Height Adjuster
Weight | 5 | 10 | 10 | 9 | 10
Volume | 3 | 10 | 10 | 7 | 10
Catheter Diameter | 6 | 10 | 10 | 10 | 10
Catheter Length | 3 | 10 | 10 | 10 | 10
Material | 3 | 10 | 10 | 10 | 10
Compatibility Time | 2 | 7 | 8 | 8 | 8
Portability | 1 | 8 | 7 | 6 | 6
Cost | 7 | 3 | 10 | 4 | 7
reaction time | 8 | 8 | 2 | 8 | 6
Pressure / flow Control Error | 8 | 8 | 2 | 8 | 6
Operational Pressure / flow Range | 8 | 8 | 6 | 2 | 8
Automation | 10 | 10 | 1 | 10 | 9
Safety | 10 | 6 | 3 | 6 | 9
Power Requirement | 3 | 8 | 10 | 8 | 8
Score |  | 619 | 443 | 564 | 635

## Slide 18: Final Controlling Decision
Automated Drainage Bag Height Adjuster
  Chosen for:
    Simplicity / cost
    Arduino compatibility
    Based on current method but adds automation
  Next steps 
    Design pulley/motor system
    Consider effects of bag movement on flow

## Slide 19: Thank you. Questions?